## Slide 1: Notes regarding the slides
The slides are valid as at January 2021 and refer to Version 3.3 of the firmware.
The slides are based on the default settings of the device.
Slides discussing settings for which admin or super user rights are necessary or useful are identified as follows:
The slides must be selected depending on the target audience.
Special rights required

## Slide 2: WipeAlyser Training
Part 1: General

## Slide 3: WipeAlyser Analyser System
WipeAlyser is an electronic analyser.
It is used to objectively
evaluate,
store and
document
rapid drug screening test results by mobile users.

## Slide 4: Using WipeAlyser with DrugWipe
WipeAlyser is compatible with DrugWipe S tests.
The device returns a qualitative result, i.e. positive/negative.
DrugWipe can still be evaluated visually.
The WipeAlyser result is authoritative.
Drug combinations and cutoffs depend on the DrugWipe type used.

## Slide 5: WipeAlyser – a drug pre-screening device
WipeAlyser is a screening test device.
Positive results must be validated by the laboratory.

## Slide 6: Temperature
Operating temperature: +5 to +40°C (same as DrugWipe)
Storage temperature: −30 to +80°C (up to 24 hours)
The temperature is monitored by the device and checked prior to measurement.

## Slide 7: Dust and moisture
WipeAlyser is protected against dust and splash water (IP 54).
WipeAlyser is resistant to splash water (printer lid closed); however, heavy rain and long-term exposure can still damage the device.
Contamination of the optical system with dust can affect the measurement result. Run the QC test at regular intervals.

## Slide 8: Resistance to impact and vibrations
Up to a certain degree, WipeAlyser is resistant to impact and vibrations*.
But please remember that WipeAlyser is a high-end optical measuring device. It should be handled with care.
The carrying bag offers additional protection.
* in accordance with IEC 60068-2-64, IEC 60068-2-27
8

## Slide 9: Device controls
Touchscreen
Start button with On/Off function
Home button to go to the main menu
Result status LED:
: negative
: positive
: Flashing = test invalid
Status LED:
: On
: Eco mode
: Flashing = charging
: Steady light = fully charged
: Flashing = hardware error

## Slide 10: Device connections
Code scanner for barcode scanning
Connector for power supply
USB port for connecting a PC or car charger adapter
SD card slot for exporting data
A Reset button is located at the bottom in case the device “freezes up.”

## Slide 11: Accessories
Carrying bag
User Guide
Power cord and power supply

## Slide 12: Accessories
USB cable for data transfer to the PC and powering the device via car charger adapter
Car charger adapter
USB stick with PC software and PDF user guides

## Slide 13: Accessories
Roll of printer paper (thermal paper, 56 mm wide)
Styluses
Quality control tests (QC tests)

## Slide 14: WipeAlyser Training
Part 2: How to perform a measurement

## Slide 15: Start-Up
Press the Start button to power on WipeAlyser.
The welcome screen appears.

## Slide 16: Status bar

## Slide 17: Login
Login requires a user name and password.
The administrator is responsible for user management.
You can also log in as standard user. This keeps the user anonymous and can be disallowed by the administrator.

## Slide 18: Self-check
After login, the device runs through a self-check.
This checks the technical condition of WipeAlyser.

## Slide 19: Self-check
Battery sufficiently charged
Sensors operational (e.g. level sensor)
Interfaces operational (e.g. USB)
GPS signal received (outdoors only)
Temperature within allowed measuring range
Time/date set and plausible
No QC tests or service due

## Slide 20: Self-check
Notification (yellow) means that measurements are still possible.
Error (red) means that action is required and the device is not ready for measurement.

## Slide 21: Service info
You can check the date of the most recent QC test and most recent maintenance under Service Info.
If a QC test or maintenance is due, the appropriate notification is highlighted in red.
The firmware version is specified below this.

## Slide 22: Service info
When using the device for the first time, the intervals are inactive and need to be activated once.
The user is prompted to do this automatically after the device has been switched on for the first time.

## Slide 23: How to prepare for a measurement
The DrugWipe type can be selected manually or by scanning the barcode on the tubular bag.

## Slide 24: 2 readout methods after sample collection
Sample collection according to DrugWipe instructions for use
The test is initiated in WipeAlyser and automatically read after a preset time.
The test is initiated outside WipeAlyser, the time is taken manually and then read in WipeAlyser.
Timer enabled
Timer disabled

## Slide 25: How to perform the analysis
While the test develops, place the device on a level surface or hold it horizontally and do not tilt.
Prolonged or excessive tilting (>30°) will abort the measurement.
If the device detects a tilt, the user is alerted by the water level symbol (and possibly an acoustic signal):
Do not move or withdraw the DrugWipe during the measurement.
Make sure that the DrugWipe is inserted all the way (otherwise, level sensor will prevent measurement).

## Slide 26: Result display
The test result on the left is positive for amphetamines/methamphetamine and opiates and negative for cannabis and cocaine.
The View image button lets you enable a schematic view of the lines on the DrugWipe.
Red (positive) test lines will flash

## Slide 27: Result display
If the result is invalid, the measurement must be repeated. One reason for invalid results could be excessive tilting of the device during the measurement.

## Slide 28: Data Input
The administrator configures the fields as mandatory or optional.
You use the scanner icon to activate the integrated scanner and enter a barcode.

## Slide 29: Data Input – Summary
A summary appears after data input.
The test subject can sign using the finger or stylus.
After storing the record, it can be output using the built-in printer.
The electronic signature is only stored in the device.

## Slide 30: Printing
The printout contains all the data fields that were filled out (except for the test subject’s signature).
It also shows all the device and test data.

## Slide 31: How to view stored results
Select Results from the main menu to view stored records.

## Slide 32: Result display
Details lets you view all the stored data.
Display image lets you view a screenshot of the DrugWipe read out window and the electronic signature.
Edit lets you retroactively modify the entries in the record. (Note: The Edit function may require administrator rights.)
The result itself is not editable.

## Slide 33: Maintenance & service
WipeAlyser must be serviced and possibly calibrated every two years.
If the device is due for service, a warning appears during start-up.
Servicing and calibration tasks are performed by the manufacturer.
Please contact your service partner.

## Slide 34: QC test
The QC test interval is configured by the administrator.
Insert the QC test cartridges one by one and confirm by selecting Test 1 or Test 2.
In case of abnormalities or implausible measurement results, we recommend running an unscheduled QC test.

## Slide 35: QC test
If the QC test is a pass, the test interval is automatically reset.
In case of an error message, the QC test must be repeated.
If the test is another fail, please contact your service partner.

## Slide 36: WipeAlyser Training
Part 3: Settings

## Slide 37: Device info/check
Device info/check lets you view device information.
If Scan mandatory is activated, the barcode on the DrugWipe tubular bag must be scanned before each test in order to verify the test type, batch and shelf life.
Special rights required

## Slide 38: Device info/check
In this menu it is possible to manually initiate a self-check or QC test.
Export lets you print or save the device information to an SD card, for example in case of a complaint.

## Slide 39: Log file
The log file is a record of all the actions performed on the device.
Select individual entries to view their details.

## Slide 40: Language
Language lets you configure language settings.

## Slide 41: Date/time
Date and time can be set manually or automatically via GPS.
Without GPS reception, the date and time can only be entered manually.
Format lets you configure the date format.

## Slide 42: Display/eco
Decreasing the display brightness will save power.
In Eco mode, the display will be dimmed while the device remains powered on (blue LED).
WipeAlyser will automatically power off after the time defined under Auto-off has elapsed.

## Slide 43: Sound
Test result signal will indicate completion of the measurement by an acoustic signal.
Signal level sensor will warn the user in case of excessive tilting of the device.
Uncritical warnings can be indicated by an acoustic signal if desired.

## Slide 44: Printer
The Printer menu lets you generate a test print.
The Paper forward and Paper backwards buttons lets you advance or retract the printer paper.
Paper feed is automatic when installing a new roll.

## Slide 45: Advanced settings
By default, advanced settings are restricted to administrators and super users.
Administrators can grant other users individual rights and create new users.
Special rights required

## Slide 46: User management
Login – With user data only lets you disable the default user.
In this case, access to the device requires the unique user ID.
Special rights required

## Slide 47: User management
User summary lets you view the currently stored users.
The administrator can define two user levels: User and Super user.
Special rights required

## Slide 48: User management
New user lets you create a new user profile.
The password to be assigned must be entered twice and can be displayed if necessary.
Rights management lets you define the user group.
Special rights required

## Slide 49: User management
Deletion of a user cannot be undone.
The data entered for the deleted user are retained.
Special rights required

## Slide 50: User management
Rights management lets you define the changes that can be made to users and super users.
Generally, administrator rights are never restricted.
Special rights required

## Slide 51: Record
Record lets you define the data entries that are mandatory or optional.
Disabled fields are greyed out during input.
Either Name or Drivers license ID or Identity card no. must be a mandatory field.
Special rights required

## Slide 52: Sensors
The GPS sensor can be deactivated completely.
The current temperature and details about the correct level can be stored with every record.
If the temperature is outside the permitted range, the system can issue a warning or the test can be disallowed.
Special rights required

## Slide 53: Sensors
The timer can be switched on or off as standard. It is also possible to make this setting switchable so that, for each test, the user can decide whether the result is to be read immediately or not until the test runtime has elapsed.
Special rights required

## Slide 54: Self-check
We recommend running the self-check each time the device is powered on.
Errors that are detected during the self-check can be defined as critical and disallow measurement.
Device errors that prevent measurement are always critical.
Special rights required

## Slide 55: Service/QC tests
Notifications that service or QC test is due can be disabled.
QC test – Interval lets you define the intervals for required QC tests.
The service interval is configured by Securetec when the product ships (typically 2 years).
Special rights required

## Slide 56: Service/QC tests
If QC test mandatory is activated, analysis cannot be performed after a failed QC test until another QC test has been successfully completed.
Special rights required

## Slide 57: Anonymise
Anonymise will irreversibly delete all personal data from the device’s internal memory.
The (anonymised) measurement data are retained.
Always anonymise stored data before the device is transferred to third parties (privacy).
Special rights required

## Slide 58: Import/Export
The stored data can be exported to an SD card from which it can later be re-imported. This is useful for backing up the device.
Exported data are encrypted and only readable by a WipeAlyser (but not necessarily the same device).
Special rights required

## Slide 59: Data structure on the SD card
…
…
…
…
…
…
Log file
Device info file
Data export (proprietary database format, can be read via WipeAlyser –> Import)
Fullimage1_0.bmp
Can be read with text editor
Any exports are stored on the SD card under:
100. measurement on 12/05/2020
Device serial no.

## Slide 60: WipeAlyser Training
Part 4: DATA MANAGER software

## Slide 61: How to install the software
The installation file is provided on the USB stick included with the delivery.
The software runs on Windows (version 7 or later).
Prior to and during installation, a WipeAlyser must be connected to the PC and powered on to load the required drivers.
Desktop icon

## Slide 62: Login
Log in as User or Administrator.
Administrators are also able to delete data from the device.
In both cases, the preset password is “securetec”.

## Slide 63: Change password
System – Change password should be used to change the password after the first login.

## Slide 64: Change password
System – Change password should be used to change the password after the first login

## Slide 65: Info
System – Info shows information about the system.

## Slide 66: How to set the language
Language lets you set the language.

## Slide 67: How to connect WipeAlyser
Click Connect to connect the WipeAlyser to the software.
After successful connection, the device appears in the window.
WipeAlyser detected

## Slide 68: Loading results
Clicking Load results generates a list with the records stored on WipeAlyser.

## Slide 69: How to select data
To select individual rows, press and hold down the CTRL key and click on the relevant rows.
If you press and hold down the SHIFT key and click on individual columns, you can deactivate these columns so that they are ignored when you save.

## Slide 70: How to transfer data
Export all or Export selection lets you transfer all or selected records to the PC.
The data are exported to the selected location as an Excel file.

## Slide 71: How to select data
Administrators can use Delete all to delete all data from the device.
Special rights required

## Slide 72: How to select data
Delete selection lets you delete selected records from the device.
Special rights required

## Slide 73: WipeAlyser Training
Part 5: Information for partner distributors

## Slide 74: Device shipments
Make sure the packaging is shock-proof.
Two copies of the UN3481 label and the shipper’s phone number must be added to the outside of the package.
MSDS for the batteries must be placed inside the package on top of the contents.
The type designation of the battery must be noted on the commercial invoice.
Observe country-specific requirements.
Contact Securetec in case of questions.

## Slide 75: Service/calibration
Partner:
Visual inspection
Cleaning
Perform QC tests
No concerns:
Return to customer
Inspection not passed: Contact Securetec
Remote maintenance session via Teamviewer as needed
Problem persists: Send to Securetec
Problem solved

## Slide 76: Firmware updates
The firmware consists of several components:
The firmware itself (at the present time, needs to be updated by Securetec via Teamviewer)
Language Table (file on SD card): contains language packs
XML files (file on SD card): contains DrugWipe info and test parameters
Service file (on SD card)
Reference images of the various blank strips (on SD card)
Contact: Peter Hoffmann
service@securetec.net

## Slide 77: What to do after detecting a bug
Notify Securetec
Take a screenshot when the bug occurs
Document the steps leading up to the bug
Firmware version of the device in question
Serial number of the device
Log file (if applicable)
Additional info that may be relevant

## Slide 78: Thank you for your time.

## Slide 79: Contact
Max Mustermann
position
Securetec Detektions-Systeme AG
Lilienthalstraße 7
85579 Neubiberg
Germany
P +49 89 203080-1651
F +49 89 203080-1652
info@securetec.net
www.securetec.net
© Securetec Detektions-Systeme AG